## Slide 1: Off-the-shelf Commercial Compact Solar FTS for CO2 and CH4 Observations for MRV
M. K. Dubey1, Z. Butterfield1, R. Lindenmaier1, H. Parker1,
J. Hedelius2, D. Wunch2, P. Wennberg2 and F. Hase3
1Los Alamos National Laboratory, Los Alamos, NM,
2Caltech, Pasadena, CA and 3KIT, Karlsruhe, Germany
NASA – CMS Project Report Bethesda, Maryland 13 Nov 2014
dubey@lanl.gov
.
TCCON
125HR
EM27/SUN
COCCON
Notes: Gas has been produced from coal and conventional mining of shallow oil and gas reservoirs in the 20th century. In the 21st century US is undergoing a transformation in oil and gas exploration by digging deeper into the gas rich shale to extract natural gas profitably to achieve energy independence. US natural gas production has increased dramatically to 23 trillion cubic feet/year in 2011 and is projected to grow to 33 trillion cubic feet/year in 2040. It has moderated the net CO2 emissions from the US that have dropped over the last few years since 2008.

## Slide 2: Outline: Enable Off-the-Shelf Validation Technologies in the Developing World
Atmospheric verification of CH4 and CO2 fluxes from ground (TCCON) and space is feasible.
TCCON (high resolution solar FTS’s to ground truth satellites) has gaps in developing regions due to its large size, high costs and operational complexity.
Harness new compact commercial low resolution solar FTS that is inexpensive to fill gaps and open difference measurements of regional fluxes.
Compared performance of the compact FTS at power plant, urban and remote TCCON site.

## Slide 3: TCCON Valuable Science in Developed Nations with Gaps in Developing Regions
Constrain global carbon fluxes to improve carbon cycle models
Primary validation (ground-truth) dataset for satellite instruments
  GOSAT, OCO-2 and others
Methane emission verification over US
Can an off-the-shelf compact and robust instrument fill gaps without compromising accuracy and long term stability?
TCCON Station
WLEF Tall Tower
OCO-2
Park Falls, Wisconsin, USA

## Slide 4: Four Corners Methane Hot Spot in Space & Time
Four Corners
LANL Four Corners TCCON (11-13)
Persistent morning CH4 high when site is downwind of source E/SE of site (70% gays)
Kort et al. GRL 2014

## Slide 5: Simulations (WRF) with CH4 inventories ( Edgar) are a factor of 3.5 lower than observations
SCIAMACHY (03-09)
TCCON (2011-2012)
Emissions from Four Corners of ~0.59 Tg CH4/yr have persisted 2003-2009, 2012. This is 10% of US EPA estimates of emissions from natural gas.
Source not related to recent shale boom with high-volume hydraulic fracturing and horizontal drilling , but likely established fossil fuel extraction activities
Kort et al. GRL 2014

## Slide 6: Evaluate Compact Commercial Low Resolution Solar FTS with Inbuilt Camtracker against TCCON to fill gaps
EM27/SUN, Pendulum Res: 0.5 cm-1 InGaAs Det. 6000-9000 cm-1 Inbuilt Camtracker 470 x 630 x 350 cm Mass: ~25 kg Cost ~ $100K Mobile
TCCON Bruker 125HR, Res: 0.014 cm-1 InGaAs Det. 3800-12000 cm-1 External tracker 20 foot seatainer ~500Kg Cost ~ $600K Stationary
TCCON
125HR
EM27
Tracker
Tracker
TCCON
EM27
Use TCCON as a standard for EM27/SUN
Reduce, size, cost, operational ease

## Slide 7: Four Corners (Power plant) Daily Variations
CO2
CH4
EM27 tracks
125HR
EM27 is high by
20 ppb

## Slide 8: Four Corners Daily Variation (Cont.)
CO2
CH4
Tight correlations demonstrate EM27 tracks 125HR changes in CO2 & CH4
EM27 CH4 is high by 20ppb
*

## Slide 9: Caltech Daily Variation (Urban)
CO2
CH4
EM27 tracks 125HR changes in
CO2 and CH4 very well
EM27 CH4 is higher that 125HR by 20 ppb

## Slide 10: Caltech Diurnal (Urban)
CO2
CH4
Tight correlations demonstrate EM27 tracks 125HR changes in CO2 and CH4
EM27 CH4 is high by 20ppb
*

## Slide 11: Los Alamos Diurnal Profile (Remote)
EM27 tracks 125HR
CO2 and CH4 well including slow variations.
Not enough variation to compare changes

## Slide 12: Conclusions and Future Work
Atmospheric remote sensing verification of CH4 fugitive leaks
Demonstrated EM27 tracks daily CO2 and CH4 changes measured by the 125HR in power plant, urban and remote well.
The EM27 is biased high by 20 ppb relative to 125HR.
The EM27 that was shipped from Germany and has been driven extensively over a year is remarkably stable.
2015: TCCON campaigns in OK, WI & possibly Eureka.
2016: Propose deployment to Brazil and Philippines
Longer: ILS monitoring and calibration protocols are being developed at KIT to establish a EM27 network (COCCON) with KIT (4), LANL (1), Caltech (1), Harvard (2) instruments.

## Slide 13: Regional Atm. Carbon-Water Coupling Observations in Amazon Rainforest to Evaluate CLM (Dubey, TES)
FTS monitors Regional Column CO2, CO, CH4, N2O, H2O/ HOD every minute
2 In situ Picarros monitoring CO2, CO, CH4 & HOD/H2O
FTS can monitor OCS & H2CO
Successful Install - Automated Operations Started 9/30/14
OCO-2 Validation – Target site programmed by NASA
Integrated with scaling CLM tasks for NGEE-Tropics
http://campaign.arm.gov/goamazon2014/documents/AmzonSolarFTS-Science-abstract.pdf

## Slide 14
Paramaribo
Manaus
Porto Velho
Volta Redonda
Manaus
TCCON
125HR
InSb (NIR)
InGaAs (MIR)
Si (vis)
Picarro
CO, CO2, CH4,
H2O, HOD
Target
Gases
CO2, CH4, CO, N2O
H2O,
HOD(ET)
OCS (Photosynth)
H2CO(BVOC)

## Slide 15: TCCON Instruments
  Ground-based Solar Fourier transform spectrometers
  Remote sensing of total columns of CO2, CH4, N2O, CO, H2O, HDO, O2 via solar absorption
  Divide trace gas columns by O2 column to get dry-air mole fractions: XCO2, XCH4, XN2O, XCO, XH2O, XHDO
Molecule | Precision | Accuracy
CO2 | ~0.8 ppm | ~0.8 ppm
CH4 | ~5 ppb | ~7 ppb
N2O | ~1.5 ppb | ~3 ppb
CO | ~0.5 ppb | ~4 ppb

## Slide 16: Portable mini-mobile FTS tested at 4-Corners, Los Alamos & Los Angeles
Pendulum FTIR
Resolution: 0.5 cm-1
470 x 630 x 350 cm
~25 kg incl. tracker
Need for portable and robust instrument Without compromising
Accuracy and long term
Stability! EM27/SUN

## Slide 18: TCCON confined to developed nations
Gaps in developing regions